The PACS managers reference manual
Problem Solve Paradigm:
1. IF a problem is caused by a user THEN
  (use the system to correct the problem) AND (educate the user)
2. ELSE IF a problem is caused by a system configuration fault
  THEN pass the problem over to those responsible for system maintenance
3. ELSE IF a problem is caused by a system fault
  THEN pass the problem on to the system developers
If followed through, the above steps will resolve all system problems.

Last modified:
Last printed:
Version no: 3.1

Title	PACS managers reference manual
Author/Owner
Created	, 22. Dec. 2016
Last modified	,
Last printed
Source	C:\Users\Peter\Downloads\14824173769051468a.odt
Configuration reference	U:\Radiology\Group\PACS documents / RJAH PACS Business continuity plan V1.2.pdf / PAS down activity work sheet.pdf / Intranet Radiology Policies / CRIS down activity work sheet.pdf / CRIS down activity work sheet V2.pdf
Date	Version	Amendment history
3 apr 2008	0.1	First draft.
20 apr 2008	1.0	First release for Paul Dallow.
15 jun 2013	3.0	Go-live date with CRIS locally hosted.
06 Sept 2013	3.01	Go-live date with Synapse PACS.
21 Sept 2013	3.02	First updates to document to incorporate Fuji system.
08 Nov 2013	3.186	Included exam code upgrade information.
11 Nov 2013	3.190	Removed more obsolete text. Added section on merge patient, RIS scheduler and door codes update.
21 Nov 2013	3.194	Update to RIS scheduler section. Added DICOM modality abbreviation list.
28 Nov 2013	3.199	Various update to outdated paragraphs.
13 Dec 2013	3.205	Added section 4.3 on creating audit report extracts.
24 Dec 2013	3.208	Added section 4.8.1 listing common admin searches.
09 Jan 2014	3.209	Added Synapse mobility user guide.
10 Jan 2014	3.210	Amended CRIS statistics output server location.
22 Jan 2014	3.211	Amended section on IEP and CD imports.
24 Jan 2014	3.214	Update to create new PACS account paragraph.
27 Jan 2014	3.216	Added note on restarting Etiam CD burner services.
14 Mar 2014	3.218	Updated notes on voice recognition.
15 Apr 2014	3.220	New section on CR retrieval of old images.
23 Apr 2014	3.221	Update to VR section.
1 May 2014	3.223	Added section on bar code reader.
8 May 2014	3.239	Added section on modality integration. Update to CRIS record merging section.
13 May 2014	3.246	Update on anomalies section.
15 May 2014	3.249	Update to the system diagram in Chapter 1.
29 May 2014	3.253	Added a section on regular maintenance searches for Synapse.
2 June 2014	3.255	Note on RIS users accounts being added to PACS automatically.
9 July 2014	3.275	Added information about CRIS urgency colouring. Added section on burning CD from CR console.
29 July 2014	3.284	Added section on Series Picker Toolbar separation bar
3 Sept 2014	3.291	Update to FTP settings regarding CRIS
4 Sept 2014	3.293	Added instruction to install Synapse 3d
13 Oct 2014	3.305	Updates to the section on PACS anomalies.
13 Oct 2014	3.305	Added recommended label printer for HSS.
23 Oct 2014	3.314	Updated outdated paragraphs on CRIS and PACS status changes.
13 Nov 2014	3.324	Updates to reflect changes received from upgrade to CRIS version 2.09.10p14.
5 Dec 2014	3.327	Added instruction to set up enhanced logging for CRIS.
11 Dec 2014	3.328	CR instructions on resolving when the wrong patient has been picked up from a CR work list.
22 Dec 2014	3.339	Update on study folder status changes on PACS.
23 Jan 2015	3.341	Added section on Study Not Reset anomalies.
24 Feb 2015	3.342	Note added to IEP section on Synapse image overwrite when existing UID is found.
27 Feb 2015	3.353	Added section on CD burning, including anynimised.
19 March 2015	3.354	Update to add new referrer section.
27 April 2015	3.360	Added reference to GE old image store
6 May 2015	3.363	Added section on text messages going out on CRIS appointments.
15 July 2015	3.373	Added section on how to opt out of text messaging.
07 Oct 2015	3.376	Added section on PAS – CRIS interface.
29 Dec 2015	3.385	Added new knowledge to 3.2 CRIS PACS DTI.
29 Mar 2016	3.392	Added detail to IEP import section.
4 April 2016	3.394	Added section anti-virus exclusion for CRIS and PACS.
9 June 2016	3.398	Updated the CRIS dictation file section.
7 July 2016	3.400	Added notes on Synapse monitor configuration.
10 Aug 2016	3.405	Updated information on CR configuration.
13 Sept 2016	3.408	Added 4.1.5.1 Image quality notes for Synapse mobility
06-10-16	3,41	Updated CRIS client ID reinstall
10-10-16	3,42	Update on 5.4 CD anonymised exports
27-10-16	3,43	Update CR user utility section Error: Reference source not found
21-12-16	3,46	Added section on D-Evo II DR system.

# RIS and PACS system overview

Fig 1.1 shows an overview of the information flow through the RIS/PACS system.
CRIS (Fuji/HSS)
Our RIS[1] (CRIS) is the main driving force. CRIS is a JAVA[2] based application and is running on Trust PC clients. Like a web page, each time CRIS is started the source is retrieved from a server off-site. This makes CRIS sensitive to network delays.
[1] RIS = Radiology Information System
[2] JAVA = Platform (computer) independent programming language, developed by SUN.
CRIS holds a copy of the demographics database found on our PAS (Lorenzo) and updates on PAS are also fed through to CRIS.

Worklists provider (Fuji)
Fuji does not feature a broker. Worklists are fetched from the Fuji DICOM server provided by the same IP as the main PACS (10.211.144.150). Sometimes modalities[3] don't like the same AET to be used for both PACS destination and broker. In that case a different AET can be used (SYNRJAHSCP2 instead of SYNRJAHSCP).
[3] Siemens Acuson US for example
Request card scanning (Trust)
Request cards are scanned via CRIS on A4 flat bed scanners and the resulting images are stored on a Trust FTP[4] server.
[4] FTP = File Transfer Protocol (network protocol to exchange files)
Voice Recognition (Dragon)
To speed up the report turnaround, voice recognitions software has been installed on as default (but not always activated) on all reporting workstations. The voice profiles are stored on a IT provided network server //opserver2/profiles/ and each user needs to have read/write to this location.
HSS CRIS has not designed a specific radiologist work flow for VR and the radiologist essentially steps in at the typing secretary in the work flow.
Synapse Quarantine server (Fuji)
The Fuji PACS is divided into two boxes; main PACS and the Quarantine server. The RIS interfaces drives the main PACS server. The Quarantine server can be used to store
studies temporarily, maintaining a foreign patient ID (non RL1 number). Studies arriving via IEP or CD are imported into the Quarantine server first. Dictation Voice Files

Fig 1.1: System diagram and information flow
Dictation voice files are stored via CRIS on the Trust FTP server. The reporting secretaries play recorded voice files back from the FTP server. A maintenance schedule is in place to delete old dictations.
The Quarantine server does behave differently at is allows duplicate patient IDs to exist without raising an anomaly. It treats duplicate patient IDs as separate patients because it is intended to be a dirty-data store (non-unique ID = on).
Synapse main (live) PACS server (Fuji)
The Fuji PACS server uses one door into the system (10.211.144.150) after which a load balancer distributes the information streams internally.
The size of the Synapse image archive is 8 TB[1].
[1] TB = Terra Byte. 1TB is 1000 GB which is 1.000.000 MB which is 1.000.000.000 kB. The average CR image is about 15MB. CT and MR studies can be up to several 100MB.
In the main PACS patient ID's have to be unique (non-unique ID = off), therefore a patient mismatch anomaly is produced and does not allow duplicate Ids. This enforces clean data.
Synapse backup/DR server (Fuji)
All studies are backed up to a DR server which is a full duplicate of the live PACS system. In case of a disaster in the live PACS server room, Fuji can switch to this system quickly.
Orthopaedic templating (Meridian Orthoview)
Orthopaedic templating is a web software package which can be launched from PACS. It creates DICOM vector drawings over the top of the DICOM images to allow surgeons to estimate size of possible implants required and/or do other pre-operative planning.
In order to use DICOM images for Orthopaedic Templating, the image needs to be calibrated (i.e. x pixels of the image represent x cm in the real world). This is done by either
1. estimating the magnification (e.g. a XR Pelvis is on average 115% magnified)
2. placing an object of known size on the x-ray and use it to calibrate the image
3. use a magnification ruler and annotate the image (method developed in the RJAH)
EPR (Trust)
The Trust has developed its own Electronic Patient Record system (EPR). A copy of the report created on CRIS is sent to EPR.

# Index